## Slide 1: LXI 1.6 - 2022 Update Differences between LXI 1.5 and LXI 1.6 - 2022

## Slide 2: LXI 1.6 in Summary
Updated Device Spec
  Incorporated the Clarification Document
  Rule to support HTTPS Web server
    Permission to still support HTTP if required
  New URL alias: <hostname>/lxi to get to LXI Welcome Page
  Enable/disable mDNS
  Permission to have a blank password is now deprecated
New Extended Functions
  Security (see LXI Security slides later on)
  API (see LXI Security slides later on)

## Slide 3: Extended Functions – no changes
No changes to these Extended functions
  Eventlog
  Event Messaging
  Timestamped Data
  Wired Trigger Bus
  VXI-11
  Clock Synchronization
    Still requires at a minimum IEEE1588-2008(v2) and not IEEE1588-2019 (v2.1)

## Slide 4: Extended Functions – minor changes
HiSLIP
  Implement IVI HiSLIP 2.0 instead of 1.1
    IVI HiSLIP 2.0 added rules for TLS and authenticated connections, but they are optional
    Therefore, LXI HiSLIP does not require them either
    If you implement the Security Extended function, then you do need to support TLS and authenticated connections

## Slide 5: Extended Functions - major Changes
IPv6
  LXI IPv6 written 10 years ago
    greater adoption of IPv6 especially in US Gov
    RFC’s have changed numerous time
    NIST maintains the NIST IPv6 Profile. A definitive list of all current RFC’s pertaining to IPv6.
  IPv6 compliance rules vs. LXI IPv6 rules
  Recommendations Static IP and DHCPv6 are now rules
  Added rule to be able to enable/disable IPv4 or IPv6
  Enable/Disable Privacy Setting now becomes a rule
  Added rule that if you support any extended functions on IPv6 then they have to work on IPv6 as well. This was a recommendation previously. E.g Clock Sync, Event messaging.

## Slide 6: Other Specifications
Guide to LXI Documentation
  List of all LXI specs with dates and versions for LXI 1.6
LXI Example and Reference Material
  The example screenshots of web pages, to show support for LXI 1.6, will be updated when the Reference Design gets released

## Slide 7: LXI 1.6 - 2023 Update Differences between LXI 1.6 - 2022 and LXI 1.6 - 2023

## Slide 8: LXI 1.6 - 2023
Minor updates to these specifications from LXI 1.6 - 2022
LXI Revision 1.6 2023 | Version | Date | Notes
Device Specification | 1.6.1 |  | Minor change
IPv6 Extended Function | 2.1 |  | Minor change
Security Extended Function | 1.1 |  | Minor change
API Extended Function | 1.1 |  | Minor change
VXI-11 Extended Function | 1.1 | 2023-06-26 | Editorial change
HiSLIP Extended Function | 1.4 |  | Minor change
LXI Example and Reference Material | 1.2 |  | Minor change
Operating Procedures | 1.93 |  | Minor change
Test Procedures | 1.6.1 |  | New

## Slide 9: LXI 1.6 - 2023
Device Specification changes
  1.4.4.2 RULE – Conformance Requirements
    List of all LXI specifications that comprise LXI 1.6 - 2023
  1.4.6 updated functional specification on Web page
    1.6 LXI Device Specification 2023
  Permission 9.8.1 – to allow blank passwords
    In 1.6 - 2022 blank passwords permission was deprecated
    Now want to make sure blank passwords are not allowed so 9.8.1 is “revoked” – not allowed

## Slide 10: LXI 1.6 – 2023 (cont1)
Additional Device Specification changes
  Rule 10.2 and 10.2.3
    Used to reference the LXI Reference and Example material document for xml identification examples.
      We have removed all the xml stuff from this document, so it now references the API specification and GitHub

## Slide 11: LXI 1.6 - 2023 (cont2)
IPv6 Extended Function - Rule 21.2.9 about individual config options for IPv6
  Now allows Auto vs Manual just like IPv4 does
    This Windows behaviour for configuring both IPv4 and IPv6
HiSLIP Extended Function
  Deprecated recommendation to add Visa Address to the mDNS text record
  Corrected Rule 20.9.2 referencing HiSLIP sub address fields - a carry over that should have been deleted before releasing 1.6 - 2022

## Slide 12: LXI 1.6 - 2023 (cont3)
LXI API Extended Function
  Miscellaneous editorial changes – see history table in this document
  Added new Identification schema (v2.0)
LXI Security Extended Functions
  Added rules regarding NIST 800-52 (TLS version and cipher suites)

## Slide 13: LXI 1.6 – 2023 (cont4)
Example and Reference Material
  Removed all XML examples as formatting in Word document issues and no XML schemas with the examples
  Put examples in GitHub where schemas are
  Maybe on LXI website as well?
Single LXI Test Procedure document
  We had separate IPv6 and HiSLIP test procedure docs in LXI 1.5
  Decided to have one “Source of truth”
  Generated by Kerberos

## Slide 14: LXI 1.6 – 2023 (cont5)
Operating Procedure
  Major release is something that is not backwards compatible or at discretion of BOD
  Minor release is adding new features and can now be deprecated and replaced with the same minor version but with something to distinguish between the two. LXI 1.6 2022 replaced with LXI 1.6 2023
Bylaws
  Strategic Members -> Strategic
  Participating Members -> Participating
  Advisory Member -> Informational Members
  New University Members

## Slide 15: LXI SharePoint / Notebook
All 1.6 specs are here
  LXI GitHub Website
    https://public.lxistandard.org/specifications.html
  LXI SharePoint:
    LXI_1.6_Specifications

## Slide 16: LXI Security Overview (with IVI) Customer Updates 2021-10-14

## Slide 17: Purpose of IVI/LXI Security
Establish response to strong customer request for security
  Including practical requirement for interoperability
Establish a security baseline that customers can reference
Setup an LXI CA to sign IDevIDs
Companies working together to understand customer needs
Security is Complex:
  213(API)+17(Security) RULES
  Device spec: 75 rules

## Slide 18: Status
LXI has completed a “stable” draft of the LXI Security and LXI API extended functions
  Will wait for completion of a trial implementation to ballot
  Doing some rework of LXI IPv6 requirements, will impact the schema
IVI has completed a first draft of VISA-C to support security
  Impacts Drivers, since they only require secure address string
IVI support for VISA.NET/VISA-COM drafts are ready for ballot

## Slide 19: LXI Security Standard Provisions
HTTPS (which will be required in LXI Device specification 1.6)
SCPI must provide a secure connection
  HiSLIP
  Raw SCPI over TLS, Telnet in scope of specification
Optional Client Authentication
Flexible Configuration via REST API
  Device configuration needed for security
  Client credentials to the device
  Manage the device credentials (certificates)

## Slide 20: LXI Authentication
Device authentication using IDevID or LDevID
  IDevID
    LXI has a CA and will sign certificates for members
    Devices may self-sign, or use other CAs
  LDevID
    Standard provisioning and management of LDevID with REST API
Client authentication using either:
  Username and Password
  Fingerprinted certificates
  Certificates authenticated based on a root certificate

## Slide 21: LXI Authorization
Authorization only provided for the API
  No concept of certain instrument users can access certain functions
  Generally, users only need to be authenticated
API KEY
  Client provides a shared secret key with every API call
Username/Password
  The client user list includes a tag of API-authorized users
  Utilizes HTTPS security scheme to pass username and password

## Slide 22: LXI Configuration API
The API is currently just for security, may be extended
Common Configuration
  Single document with:
    Device configuration
    Client credentials
  Goal to permit same Common Configuration for all devices in a system
Device-specific/Automatic Configuration
  Basic network parameters (DNS,DHCP, addresses, SLAAC, etc)
General read/write support
  Get capability information on a read

## Slide 23: Protocol Configuration
<Interface name="lxi" enabled="true" insecureMode="true" LXIConformant=""
otherInsecureProtocolsEnabled="true">
<Network>
<IPv4 enabled="true" DHCPEnabled="true" autoIPEnabled="true" dynamicDNSEnabled="true" />
<IPv6 enabled="true" DHCPEnabled="true" dynamicDNSEnabled="true"
mDNSEnabled="true" pingEnabled="true" prefix="123"
privacyModeEnabled="true" SLAACEnabled="true" />
</Network>
<HTTP operation="enable" port="80" />
<HTTPS clientAuthenticationRequired="true" port="443">
<Basic enabled="false" />
<Digest enabled="true"/>
<Service name="LXIAPI" enabled="true" />
</HTTPS>
<SCPIRaw enabled="true" port="5025" />
<Telnet enabled="true" clientAuthenticationRequired="false" port="5024" TLSRequired="true"/>
<SCPITLS port="5026" capability="99" clientAuthenticationRequired="true" enabled="true" />
<HiSLIP enabled="true" port="4880" encryptionMandatory="true" mustStartEncrypted="true">
<ClientAuthenticationMechanisms>
<PLAIN enabled="true"></PLAIN>
<SCRAM enabled="true"></SCRAM>
<MTLS enabled="true"></MTLS>
</ClientAuthenticationMechanisms>
</HiSLIP>
<VXI11 enabled="true" />
</Interface>

## Slide 24: Client Credentials
<ClientAuthentication>
<ClientCredential APIAccess="true" password="secret" user="Suzie"/>
<ClientCredential APIAccess="false" password="supersecret" user="Billy"/>
<ClientCertAuthentication>
<RootCertPEM>
-----BEGIN CERTIFICATE-----
MIIEczCCA1ugAwIBAgIBADANBgkqhkiG9w0BAQQFAD..AkGA1UEBhMCR0Ix
EzARBgNVBAgTClNvbWUtU3RhdGUxFDASBgNVBAoTC0..0EgTHRkMTcwNQYD
VQQLEy5DbGFzcyAxIFB1YmxpYyBQcmltYXJ5IENlcn..XRpb24gQXV0aG9y
aXR5MRQwEgYDVQQDEwtCZXN0IENBIEx0ZDAeFw0wMD..TUwMTZaFw0wMTAy
MDQxOTUwMTZaMIGHMQswCQYDVQQGEwJHQjETMBEGA1..29tZS1TdGF0ZTEU
MBIGA1UEChMLQmVzdCBDQSBMdGQxNzA1BgNVBAsTLk..DEgUHVibGljIFBy
aW1hcnkgQ2VydGlmaWNhdGlvbiBBdXRob3JpdHkxFD..AMTC0Jlc3QgQ0Eg
THRkMIIBIjANBgkqhkiG9w0BAQEFAAOCAQ8AMIIBCg..Tz2mr7SZiAMfQyu
AQH/MA0GCSqGSIb3DQEBBAUAA4IBAQC1uYBcsSncwA..DCsQer772C2ucpX
xQUE/C0pWWm6gDkwd5D0DSMDJRqV/weoZ4wC6B73f5..bLhGYHaXJeSD6Kr
XcoOwLdSaGmJYslLKZB3ZIDEp0wYTGhgteb6JFiTtn..sf2xdrYfPCiIB7g
BMAV7Gzdc4VspS6ljrAhbiiawdBiQlQmsBeFz9JkF4..b3l8BoGN+qMa56Y
It8una2gY4l2O//on88r5IWJlm1L0oA8e4fR2yrBHX..adsGeFKkyNrwGi/
7vQMfXdGsRrXNGRGnX+vWDZ3/zWI0joDtCkNnqEpVn..HoX
-----END CERTIFICATE-----
</RootCertPEM>
<CertThumbprint hash="SHA-256" thumbPrint="4281B2V490b84a129weff08908fsdjfksljf" />
</ClientCertAuthentication>
</ClientAuthentication>

## Slide 25: LXI Certificate Management API
Several conventional REST APIs to:
  List certs
  Delete a cert
  Get a cert
  Get a Certificate Signing Request, then Put signed certificate
  Request the device generate a self-signed certificate
No LXI requirement for SCEP or EST
  EST seems the ‘proper’ well-written secure standard, but limited support in IT infrastructures
  SCEP is commonly used, but implementations are a little divergent
  Configuration, debug, and scalability all concerns
  The API makes it practical for a client to integrate into any system

## Slide 26: IVI Address for Drivers and VISA
The address string has decoration of “Credential Information”
  TCPIP0::DevCredentials@192.168.1.0::INSTR
  Identifier includes how VISA:
    Authenticates itself to the device
    Authenticates the device
  Vendor-specific configuration completed outside of VISA
Optional syntax permits credentials directly in string (vendor specific)
  Contentious, some vendors refuse to implement (security concerns)
Why do it this way?
  Vendors utilize their own scheme to secure the credential information
  Vendors create their own applications software to collect and store credentials (difficult for consortium)

## Slide 27: IVI Extensions for Security
VISA-C with Address string updates, completed early 2021
  Additional APIs to acquire device certificate and various fields
  Address strings per previous slide
VISA-COM, VISA.NET, minor changes, ready for ballot
No changes identified for drivers – only address strings needed

## Slide 28: Future Priorities (Need Input!)
Certificate Enrollment using standard protocols
  Do we use EST, SCEP or something else?
  What does a customer need to debug and configure?
Port-based network access control
  IEEE 802.1X and subsequent standards
  How common is this in T&M systems? Is MAC registration adequate?
More complete User Authorization solution
  What users are allowed to do what?
Participation in federated/SSO systems for user Authentication and Authorization
  What systems to support? (MSFT Active Directory)
  How to use information?